Background

The FeedRite Feeding Tube is the next generation in naso-duodenal/naso-jejunal (ND/NJ) feeding. Currently, ND/NJ feeding tubes are placed with one of two main methods. The first method, fluoroscopy, accurately places tubes in the duodenum and jejunum. However, it is a notoriously slow process and often requires repeated fluoroscopic procedures, exposing the patient to unnecessary amounts of radiation. The second method relies on RF communication between the tube and a receiver placed on the xiphoid process. This method is relatively safe, but often results in improper tube placement.

Instead of using external indicators like those previously mentioned, the FeedRite Feeding Tube utilizes the body’s natural physiology to determine tube location in the gastrointestinal tract. By relying on known pH and pressure differences between the stomach, duodenum, and jejunum, FeedRite plans to accurately, safely, and efficiently place ND/NJ tubes.

Achievements since last report

We have received our feeding tube and have contacted CELA to set up a time to come test and measure our tube. The tube is more narrow than the camera and scope at the testing center, so we will not be able to insert the camera into the tube in order to mark the approximate location of the landmarks of interest (specifically, the pyloric sphincter). Instead, we will mark on the scope where along the path the entrance to the stomach and duodenum occur and use this measurement to mark the tube.

Though we were unable to purchase the specific solutions we intend to use for our testing chamber, we were able to test the pH probe using common household objects. Below is a table of the reported pH values and measured values. We will use this information to test the efficacy of our testing chamber.

| Product | Reported pH | Average Measured pH | Error |
| --- | --- | --- | --- |
| Dawn Ultra, Hand Renewal, Dishwashing Liquid | 9 | 8.32 | 7.56% |
| Clorox Toilet Bowl Cleaner with Bleach | 12.5-13.5 | 11.68 | 6.56%-13.5% |
| Publix White Distilled Vinegar (5% Acidity) | 2.4 | 2.12 | 11.7% |
| ACT Restoring Mouthwash with Fluoride | 6.6 | 6.13 | 7.12% |

We have also designed a casing that will hold the components of our system in a light portable way. We will be taking our design to be 3D printed in the near future. We hope that this improvement will drastically improve the usability of our feeding tube placement system.

While we were not able to get into contact with the company producing the pressure sensor in regards to its waterproofness, after consulting the Internet, we have determined that the entire board and sensor are waterproof besides the wire leads. Because of this, we will have to buy epoxy resin and put a few layers of it over the wire leads before we are able to test the sensor with our solutions.

Lastly, we have investigated a couple models that can be used to create a user-friendly user interface. One involves the use of Matlab while the other involves commands native to Arduino.

Problems that have arisen

The Pcard we plan to use for solution purchase requires a previously authorized form.  We did not expect the requirement of prior authorization, so we are still waiting for the approval of Pcard to purchase our anticipated liquids.

We expect the arrival of PVC pipe on Tuesday so that we can calibrate the pressure sensor by inserting it into the pipe filled with water.  However, if we could not obtain the PVC pipe by next week, we might need to use another method for calibration.  So far, our second plan is to put an object with known mass on top of the pressure sensor.  The pressure sensor will act as a force sensor.  By calculating the weight of each object, we will know the expected pressure and generate a calibration curve.

Future Steps

After the department approves our request for a P-Card, we will go to a grocery store and purchase the solutions as soon as possible.  We expect to receive the PVC pipe on Tuesday, so we will start to assemble our testing chamber and test the best combination of testing liquids.

Once we confirm the waterproofing of the pressure sensor, we will calibrate it by using either the PVC pipe or known weights.

In addition, we want to meet with our sponsor Dr. Abumrad to present him some available microsensors.  Because the limitations of time, budget and building materials, we will not be able to scale down the sensors and the entire system, but we still want to list some potential microsensors in the market which could be used for future design and modifications.

We will be incorporating additional design components that will improve the portability and usability of our device. Our ideas include an LED Arduino Screen that will eliminate the need for a computer, a user-friendly interface that will plot the pH and pressure if connected to a computer, and a casing to enclose all of the circuit components.

Finally, we will start to design our poster for design day.  We will begin by outlining the general sections we want to include in our poster.

Assessment of Schedule, Budget, and Objectives

Once we show that our design is functional, we will be able to begin the process of making the device more user friendly. In the upcoming weeks, we hope to solder the wiring of our circuit and encase all of the components in a 3D printed case. This will help us market our device as highly portable, one of the main objectives of our design. Most of the additional components we hope to incorporate into our design will not have any cost, so our budget has remained under control. An additional component that will drastically improve our design will be a more user-friendly user interface that will plot the measured pH and pressure continuously to allow the user to better identify the placement of the tip in the gastrointestinal system.

The major holdup that is preventing further progress is the time we have been waiting for the PVC pipe from Home Depot. Because this order had to be placed online, we anticipated having to wait for shipping. However, when we went to check on the progress after waiting for two weeks, we were informed that the order was lost. We placed a new order with expedited shipping and anticipate the pipe’s arrival on Tuesday of this week. After its arrival, we should be able to continue with our testing chamber and calibration of the pressure sensor. If we still have not received the pipe this week, we will calibrate the sensor using another method.

Gantt Chart