## Slide 1: The (Matthew Walker) Texas Rangers III Developing a Ventriculoperitoneal Shunt Failure Monitoring Approach for Pediatric Hydrocephalic Patients
Oral Report #4
Wednesday, March 15
Zoha Malik, Alvin Mukalel, Cole Pickney, Sungho (John) Suh, Colin Sweeney

## Slide 2: Hydrocephalus
1/1000 live births in the US
Cerebrospinal fluid drains in the brain at 600-700 ml/day (0.2-0.7ml/min)
CSF drainage from the brain is blocked
Leads to build up of CSF:
  Ventricles enlarge
  Brain swells up
  Intracranial pressure increases
    Critical at +5mmHg or more
- Loss of vision, headaches, neurological damage or death
Image source: http://r3.emsworld.com/files/base/image/EMSR/2013/09/16x9/1280x720/hydrocephalus-fnl_11140875.jpg

## Slide 3: Current Solution: Ventriculoperitoneal Shunts
Gold standard for 50+ years
Shunt surgically inserted:
  Proximally: into third ventricle
  Distally: Into peritoneal cavity
However…
  50% shunts fail within 2 years of implantation
Image source: http://umm.edu/~/media/ADAM/Images/en/12726.ashx
No means of detecting shunt failure:
  Symptoms & damage before shunt replacement
  Unnecessary replacement surgeries
    Proximal failure -> DANGEROUS
    Distal failure

## Slide 4: Project Objective Statement The Need
Measure the differential pressure between the shunt and the brain for proximal failure detection:
  If there is no difference: there is no proximal failure
  If there is a difference, there is proximal failure
    Optional distal pressure measuring component for distal failure detection
Remotely communicate when the shunt fails
The solution:
Must detect shunt failure before patient develops neurological symptoms; specifically proximal
Must remotely communicate failure with minimal error
ΔP
Communicate

## Slide 5: “Smart Shunt” Design
Notes: For cranial applications, the subQ coil preferably is positioned between the skin and the bone that forms the cranium. Electromagnetic (or inductive) coupling between external coil and SubQ coil provides for the transfer of power and data signals between the coils.

## Slide 6: Current Iteration of Design Components
Notes: Explain why fiber optics are out
Explain how strain gauges interfaced with NFC chips are feasible

## Slide 7: Strain Gauge: PDMS Window
Pore Diameter: .85 mm
# Pores: 24
# Removed Pores: 4
Drainage Surface Area: 64.1 mm2
% Reduction of Drainage Surface Area After Modification: 16.7%
2.70 mm
.625 mm
2.70 mm
1.35 mm
7.5 mm long
Strain Gauge
Proximal End
Strain Gauge
Distal End

## Slide 8: PDMS Window CAD Rendering

## Slide 9: Will Covering Pores Compromise CSF Flow?
24 Pores
12 Pores
3 Pores
18 Pores
# Pores Open | % Area Remaining | Relative Flow Rate
24 | 100% | 100%
18 | 75% | 98%
12 | 50% | 86%
3 | 13% | 75%
0 | 0% | 0%
Figure 1: Cell attachment under physiological flow at each pore
Source: Harris et al. (Modeling Shunt Failure by Tissue Obstruction)
Flow Experiment Setup

## Slide 10: RFID/NFC Communication
Physician-mediated remote sensing
Δ Pressure Sensing Mechanism: Strain Gauge
NFC Chip / RFID
RFID (radio frequency identification)uses electromagnetic waves to exchange data between an interrogator and the transponder
  Transponder is generally composed of an integrated circuit for storing and processing the data signal and an antenna for receiving power and transmitting the signal
  The electromagnetic field received from the interrogator is the only source of power to the passive transponder
NFC (near field communication) is a subset of RFID
  Allows for peer-to-peer communication
Interrogator
Transponder
Integrated Circuitry
Digital Signal Processing& Display
Demodulation
Skin
Power
Signal
ADC
Signal Conditioning
Catheter
CSF Flow
Pressure Sensor
fr = fr
Notes: EZ430-TMS37157
When a tag is placed within the alternating magnetic field created by the reader, it draws energy from the magnetic field. This additional power consumption can be measured remotely as a voltage perturbation at the internal impedance of the reader antenna. The periodic switching on and off of a load resistance at the tag therefore effects voltage changes at the reader’s antenna and thus has the effect of an amplitude modulation of the antenna voltage by the remote tag. If the switching on and off of the load resistor is controlled by the tag’s stored data stream, then this data is transferred from the tag to the reader.

## Slide 11: Arduino Implementation
Pressure Sensor
Notes: https://www.allaboutcircuits.com/projects/read-and-write-on-nfc-tags-with-an-arduino/
Red marker indicates whether or not Arduino itself is capable of sending data rather than NFC depending on our goals. Could have embedded nfc chip at a different location, while arduino is housed in a bigger space like the abdominal cavity.
Qduino (bottom right) can be used instead of read/write depending on processing power needed from the sensor.

## Slide 12: Arduino Strain Gauge Setup
INA125P instrumentation amplifier
  Gain ranges from 4-10,000 depending 6-60kΩ
Strain gauge measurement using analog in
Use Arduino to power bridge

## Slide 13: Arduino NFC Setup
Using SPI
NFC shield uses digital in/out pins
Write strain gauge info to NFC tag
Read NFC tag with a phone
Notes: 4050 level shifter needed to use the SDA/SCL pins of the nfc shield with 3.3V

## Slide 14: Construction and Testing Plan
During soldering, tape the strain gauge to a piece of aluminum to draw heat away from sensitive components
  Soldering has to be performed under a microscope and in good lighting due to small scale
Two options for wire
  Solid Cu: cheaper, less flexible
  Teflon coated Cu: flexible, more expensive, more suitable for true application
2.3 mm
6.1 mm
Vishay Strain Gauge
Notes: There are a few different particulars to address when it comes to testing

## Slide 15: Construction and Testing Plan
Develop a protocol for removal of portion of shunt
  Test cutting and PDMS blocks sealed with bonding adhesive on excess distal tubing
Final ~3 cm of proximal tubing is portion in ventricle
  Experiment with window placement in that 3 cm
  Reference: window is 7.5 mm in length, so many available positions
MR Image of Ventricles
Notes: There are a few different particulars to address when it comes to testing

## Slide 16: Further Material Needs
336-FTE: 0.127 mm diameter, silver plated Cu wire with Teflon insulation
MR1-350-130: bridge completion module
  Provides a precision 350Ω half bridge and a 350Ω dummy gage
  Output on the scale of microvolts -> need differential op amp to scale the signal up to range where Arduino can resolve
Additional PDMS and Curing agent – available from student lab
Tin-Silver Solder
Portable Strain Gauge Welding & Soldering Unit
Additional Strain Gauges
Proximal catheters from VUMC (hopefully)
MR1-350-130 Bridge Completion Module

## Slide 17
2016
2017
Today
Nov
Dec
Jan
Feb
Mar
Apr
2016
2017
Semester 1 Ends
12/19/2016
Semester 2 Begins
1/9/2017
Design Day
4/24/2017
7 days
Dec 9 - Dec 19
Finalize Semester 1 Work
15 days
Dec 19 - Jan 8
Winter Break
15 days
Jan 8 - Jan 27
Determine pressure sensing mechanism
11 days
Jan 25 - Feb 8
Construction of physical model
11 days
Feb 8 - Feb 22
Testing of pressure sensing in physical model
10 days
Feb 22 - Mar 7
Revisions of design
16 days
Mar 7 - Mar 28
Long term testing until shunt failure
20 days
Mar 28 - Apr 24
Final Revisions
12 days
Apr 7 - Apr 24
Preparation of design day deliverables
Gantt Chart