## Slide 1: The (Matthew Walker) Texas Rangers III Developing a Ventriculoperitoneal Shunt Failure Monitoring Approach for Pediatric Hydrocephalic Patients
Oral Report #3
Wednesday, February 21
Zoha Malik, Alvin Mukalel, Cole Pickney, Sungho (John) Suh, Colin Sweeney

## Slide 2: Hydrocephalus
1/1000 live births in the US
Cerebrospinal fluid drains in the brain at 600-700 ml/day (0.2-0.7ml/min)
CSF drainage from the brain is blocked
Leads to build up of CSF:
  Ventricles enlarge
  Brain swells up
  Intracranial pressure increases
    Critical at +5mmHg or more
- Loss of vision, headaches, neurological damage or death
Image source: http://r3.emsworld.com/files/base/image/EMSR/2013/09/16x9/1280x720/hydrocephalus-fnl_11140875.jpg

## Slide 3: Current Solution: Ventriculoperitoneal Shunts
Gold standard for 50+ years
Shunt surgically inserted:
  Proximally: into third ventricle
  Distally: Into peritoneal cavity
However…
  50% shunts fail within 2 years of implantation
Image source: http://umm.edu/~/media/ADAM/Images/en/12726.ashx
No means of detecting shunt failure:
  Symptoms & damage before shunt replacement
  Unnecessary replacement surgeries
    Proximal failure
    Distal failure

## Slide 4: Project Objective Statement The Need
Measure the differential pressure between the shunt and the brain for proximal failure detection:
  If there is no difference: there is no proximal failure
  If there is a difference, there is proximal failure
Optional distal pressure measuring component for distal failure detection
Remotely communicate when the shunt fails
The solution:
Must detect shunt failure before patient develops neurological symptoms
Must remotely communicate failure with minimal error
ΔP
Communicate

## Slide 5: “Smart Shunt” Design
Notes: For cranial applications, the subQ coil preferably is positioned between the skin and the bone that forms the cranium. Electromagnetic (or inductive) coupling between external coil and SubQ coil provides for the transfer of power and data signals between the coils.

## Slide 6: Current Iteration of Design Components
Issues:
Transduction method highly complex – beyond scope of our capabilities
External device is large and requires large amount of power – does not fulfill our need for monitoring in non-clinical setting
Notes: Explain why fiber optics are out
Explain how strain gauges interfaced with NFC chips are feasible

## Slide 7: Radio Frequency Identification (RFID) Communication
RFID uses electromagnetic waves to exchange data between an interrogator and the transponder
  Transponder is generally composed of an integrated circuit for storing and processing the data signal and an antenna for receiving power and transmitting the signal
  The electromagnetic field received from the interrogator is the only source of power to the passive transponder
Interrogator
Transponder
Integrated Circuitry
Digital Signal Processing& Display
Demodulation
Skin
Power
Signal
ADC
Signal Conditioning
Catheter
CSF Flow
Pressure Sensor
fr = fr
Notes: EZ430-TMS37157
When a tag is placed within the alternating magnetic field created by the reader, it draws energy from the magnetic field. This additional power consumption can be measured remotely as a voltage perturbation at the internal impedance of the reader antenna. The periodic switching on and off of a load resistance at the tag therefore effects voltage changes at the reader’s antenna and thus has the effect of an amplitude modulation of the antenna voltage by the remote tag. If the switching on and off of the load resistor is controlled by the tag’s stored data stream, then this data is transferred from the tag to the reader.

## Slide 8: RFID Communication
Telemetric Intracranial Pressure Monitoring with the Raumedic Neurovent P-tel
Source: Raumedic AG
Benefits:
RFID technology currently being used in FDA-approved continuous ICP monitor
No need for on-board power supply
  No potential leakage of battery components
  Increased failure detection lifetime
At-home recording possible
If multiple sensors are employed, modules can be configured to transmit information sequentially to ensure signal trains do not interfere
Proposed Multi-Sensor unit with multiple RFID sensing points.
Source: Patent US20090204019 A1.

## Slide 9: Arduino Implementation
Pressure Sensor
Notes: https://www.allaboutcircuits.com/projects/read-and-write-on-nfc-tags-with-an-arduino/
Red marker indicates whether or not Arduino itself is capable of sending data rather than NFC depending on our goals. Could have embedded nfc chip at a different location, while arduino is housed in a bigger space like the abdominal cavity.
Qduino (bottom right) can be used instead of read/write depending on processing power needed from the sensor.

## Slide 10: eZ430-TMS37157
PC
RFID Reader
eZ430-TMS37157
Pressure Sensor
Notes: For ultra-low-power consumption, the TMS37157 offers the possibility to completely switch off the microcontroller (if powered from VBATI) resulting in an overall power consumption of typically 60 nA.

## Slide 11
Strain Gauges: PDMS Window
Pore Diameter: .85 mm
# Pores: 24
# Removed Pores: 4
Drainage Surface Area: 64.1 mm2
% Reduction of Drainage Surface Area After Modification: 16.7%
1.45 mm
2.70 mm
.625 mm
2.70 mm
1.0 mm
7.5 mm long
Strain Gauge
Proximal End

## Slide 12: Future Directions
Separate into development teams
Order microstrain gauge
Order Arduino + decide on NFC tag
App Development
PDMS + Strain Gauge Mechanical Testing
  Fatigue
  Stress-Strain Curve
2.3 mm
6.1 mm

## Slide 13
2016
2017
Today
Nov
Dec
Jan
Feb
Mar
Apr
2016
2017
Semester 1 Ends
12/19/2016
Semester 2 Begins
1/9/2017
Design Day
4/24/2017
7 days
Dec 9 - Dec 19
Finalize Semester 1 Work
15 days
Dec 19 - Jan 8
Winter Break
15 days
Jan 8 - Jan 27
Determine pressure sensing mechanism
11 days
Jan 25 - Feb 8
Construction of physical model
11 days
Feb 8 - Feb 22
Testing of pressure sensing in physical model
10 days
Feb 22 - Mar 7
Revisions of design
16 days
Mar 7 - Mar 28
Long term testing until shunt failure
20 days
Mar 28 - Apr 24
Final Revisions
12 days
Apr 7 - Apr 24
Preparation of design day deliverables
Gantt Chart